FTIR experiment – May 1st 2017

Tasks:

- Learn to use the machine and software by doing some basic scans

- Get 5 scans of water, bacteria and bacteria diluted 1:100.

- Get scan of dried surface after water and after bacteria

Part 1

Background:

Water 10:33:

Note - The noise is only in the “dead” regions of the background - possible this is just machine noise…

Bacteria @1:100 vs water

After cleaning with EtOH - water

Bacteria 1:1

Experiment 1 - 5 scans for statistics – Bacteria in Solution

Protocol:

- Clean twice with EtOH and wipe with cotton wool.

- Sample size: 600uL, spread evenly.

- Surface temp 38C	Room temp 23C (aircon)

- Purge delay 20s

- 32 scans

- Gain = 8

- Door closed. Alone.

Note – scan takes 1:15m

Scans:

5 x Water

5 x Bacteria (~108 CFUs/ml)

5 x Bacteria diluted 1:100

11:14 - cleaned with EtOH and took Water Background

Finished scanning: 12:14

Note – scan Water 2 and Water 2* were strange and taken off analysis.

PLS diagnostic:

Done with JMP. PLS of 14 samples.

Validation: Leave-One-Out

Factors: 8

Blue – water

Yellow – diluted bacteria (defined as 0.8 for JMP)

Red - Bacteria

Experiment 2 - Scan dry

Changed the surface temp to 60C.

Cleaned with ethanol and let evaporate.

Evaporated water as background. Evaporation takes 25-35 minutes.